## Slide 1: PRES test runs: 2019 results 2021+ preparations

## Slide 2: Requirements
Beam monitoring system
Beam counter
  Number of beam electrons with uncertainty of 0.1%
  Beam intensity up to 2MHz
  Minimum beam distortions
Beam position monitor
  Position uncertainty 100um
  Feedback signal to accelerator

## Slide 3: Requirements
Beam monitoring system
Beam counter
  Number of beam electrons with uncertainty of 0.1%
  Beam intensity up to 2MHz
  Minimum beam distortions
Beam position monitor
  Position uncertainty 100um
  Feedback signal to accelerator
Scintillation counter
  absolute measurement
  pileup influence
Ionization chamber
  no pileup
  relative measurement

## Slide 4: Beam counter setup (2019)
2018 problems:
No counters readout (scope)
IC current background (~10%)
IC leaks (bad measurements)
Required precision: 0.1%
2019 Goals:
IC calibration better than 0.1%
IC background less than 0.1%
SC pair time evaluation

## Slide 5: Scintillation counters

## Slide 6: Ionization chamber

## Slide 7: Ionization chamber
Pressure
AMP Temp.
High voltage
IC current amplifier
Keithley 2000
Keithley 2700
IC Temp.

## Slide 8: Beam setup
TPC
IC
HV control
Slow control
S2,S3 position

## Slide 9: IC. Pressure and temperature
0.1% per day

## Slide 10: IC. Dark current
Full scale: 40nA
Background noise: 0.1 pA
SNR: 2.5E-6
Notes: Amplier signal -> Keithley 2000 multimeter, the same as in 2018,
but 200ms measurement period and with averaging by 10 samples.
Thus the sampling time was 2 seconds instead of 20ms in 2018.
Total measurement time - 6 seconds.
All other parameters were measured with Keithley 2700 withing the same 6 seconds during the ionization current acquisition.

## Slide 11: IC calibration
Hi->Lo
Lo->Hi
Combined
0.01%
Notes: All 4 channels!

## Slide 12: Time structure of the beam (S1 CH2)

## Slide 13: Time structure of the beam (S2 CH1)

## Slide 14: Pair time investigation
Signal width = PMT width + NINO stretching
Notes: PMT = 5ns
We add to this data calibration measurements at low intensity and evaluate tau from quadratic fit

## Slide 15
Notes: I_LIN = 20nA

## Slide 16
IC measurements
at high intensity needed!

## Slide 17: Conclusion
Background of ionization chamber current measurement: 2.5E-6
IC Pressure stability: 0.1% per day.
Ionization chamber calibrated to 0.01%
Pair time evaluation from signal stretching works, but underestimates the result (IC non-linearity?)
Difference between SC channels?
Off-axis counter can be used for indirect beam intensity measurements, but with stable beam positioning
2018 problems:
No counters readout (scope)
IC current background (~10%)
IC leaks (bad measurements)
Required precision: 0.1%
2019 Goals:
IC calibration better than 0.1%
IC background less than 0.1%
Sc pair time evaluation

## Slide 18: Requirements
Beam monitoring system
Beam counter
  Number of beam electrons with uncertainty of 0.1%
  Beam intensity up to 2MHz
  Minimum beam distortions
Beam position monitor
  Position uncertainty 100um
  Feedback signal to accelerator
Detector type: proportional chamber
Current mode: fA amplifier
Time constant: seconds
Signal processing choice
Accuracy and sensitivity investigation

## Slide 19: Beam profile MC
20 bar
5 bar

## Slide 20: Beam position monitor
Aperture: 40x40mm
Ar+30%CO2
1.8kV
Notes: 40x40mm. Wires: 30um x 2mm, cathode: Cu foil 50um, gap=2.5mm
Ar+30%CO2. 1.8kV

## Slide 21: Test rig
fA Amplifier
(8 channels)
Chamber
Source,
XY-table
ADC (32ch 16 bit)
Beam profile

## Slide 22: Signal processing
Raw signal
Gauss fit
Mean value
Parabola
Notes: Gauss, mean: sigma=3.2um
Parabola: sigma=3um (analytic fit in three peak points)
In terms of sensitivity all three methods are fine

## Slide 23: Calibration results [beam RMS=3.3mm]
Notes: Gauss – best precision
Mean – the same relative accuracy, but different slope
Parabola – sensitive to the chamber structure (wires)

## Slide 24: Calibration results [beam RMS=3.3mm]

## Slide 25: Sensitivity check [beam RMS=3.3mm]
Notes: All three methods have good sensitivity

## Slide 26: Calibration results [beam RMS=4.8mm]
Notes: For wide beam profile gauss method becomes worse due to side effects
Mean value: good sensitivity, different slope

## Slide 27: Sensitivity results [beam RMS=4.8mm]

## Slide 28: Time dependence
3s integration RC filter

## Slide 29: Time dependence
5um drift
20um ripple

## Slide 30: High voltage influence

## Slide 31: Full scale setup
4 plates x
8 channels

## Slide 32
X1
X2
Y1
Y2

## Slide 33: Conclusion
Absolute accuracy: ~70um in ±3mm range
Relative sensitivity: <20um
Gauss fit – the best for absolute position calculation
Mean value calculation – the best for feedback signal (easy to program in PLC)
Analog feedback signal: 0..10V 16-bit
Notes: Parabola – high differential non-linearity due to processing only 3 points out of 8

## Slide 34
THANK YOU!